## Slide 1: Potential Exposure To Nanomaterials
Dustiness
Bulk Density - Morphology -electrostatic forces
Dry powder liquid
Type of process
High Energy Process – Open Process
Quantity / Duration/ Frequency

## Slide 2: Safety Through the Life Cycle of a Nanomaterial
Synthesis
Characterization and Purification
Application and Material Testing

## Slide 3
State of the nanomaterial | Activity | Potential Exposure Source | Recommended controls
Bound or fixed nanostructures (polymer matrix) | ■ Mechanical grinding, alloying, etching, lithography, erosion, mechanical abrasion, grinding, sanding, drilling, heating, cooling | ■ Nanomaterials may be released during grinding, drilling, and sanding. Heating or cooling may damage the matrix, allowing release of nanomaterial. | ■ Local exhaust ventilation ■ Laboratory chemical hood (with HEPA filtered exhaust) ■ HEPA-filtered exhausted enclosure (glovebox) ■ Biological safety cabinet class II type A1, A2, vented via thimble connection, or B1 or B2

## Slide 4
State of the nanomaterial | Activity | Potential Exposure Source | Recommended controls
Liquid suspension, liquid dispersion | ■ Synthesis methods: chemical precipitation, chemical deposition, colloidal, electrodeposition crystallization, laser ablation (in liquid) ■ Pouring and mixing of liquid containing nanomaterials ■ Sonication ■ Spraying ■ Spray drying | ■ Exposures may result from aerosolization of nanoparticles during sonication or spraying, equipment cleaning and maintenance, spills, or product recovery (dry powders). | ■ Laboratory chemical hood (with HEPA filtered exhaust) ■ HEPA-filtered exhausted enclosure (glovebox) ■ Biological safety cabinet class II type A1, A2, vented via thimble connection, or B1 or B

## Slide 5
State of the nanomaterial | Activity | Potential Exposure Source | Recommended controls
Dry dispersible nanomaterials and agglomerates | Collection of material (after synthesis), material transfers, weighing of dry powders, mixing of dry powders | Exposures may occur during any dry powder handling activity or product recovery. | ■ Laboratory chemical hood with HEPA filtered exhaust ■ HEPA-filtered exhausted enclosure (glovebox) ■ Biological safety cabinet class II, B1 or B2

## Slide 6
State of the nanomaterial | Activity | Potential Exposure Source | Recommended controls
Nanoaerosols and gas phase synthesis (on substrate) | Vapor deposition, vapor condensation, rapid solidification, aerosol techniques, gas phase agglomeration, inert gas condensation (flame pyrolysis, high temperature evaporation), or spraying | Exposures may occur with direct leakage from the reactor, product recovery, processing and packaging of dry powder, equipment cleaning, and maintenance. | ■ Glovebox or other sealed enclosure with HEPA-filtered exhaust ■ Appropriate equipment for monitoring toxic gas (e.g., CO)